## Slide 1: Sensors
By Finn Haugen
(finn.haugen@usn.no)
PEF3006 Process Control Fall 2017

## Slide 2: Current loop

## Slide 3: Purpose: Getting V from mA

## Slide 4: Where is the sensor in the control loop?
Temperature
Pressure
Level
Flow

## Slide 5: Temperature

## Slide 6: Temperature sensor: Thermocouple (TC)
[http://en.wikipedia.org/wiki/Thermocouple]
A thermocouple consists of two conductors of different alloys. Whenever there is a temperature difference between the hot junction (measuring point) and the cold junction there is a voltage difference between the conductors at the cold junction which can be measured (the Seebeck-effekten).
Cold junction temperature is typically room temperature and it is measured with a a thermistor inside the measurement transmitter.

## Slide 7: Thermocouples cont.
[Autek AS]
[Autek AS]
Different types identified with letter codes: K, E, J, N, etc,
They have different voltage/temperature sensitivity and different environmental properties. Tables can readily be found on the Internet.
Type K is most common:
Cromel + Alumel (both are nickel alloys).
Sensitivity: 41 µV/°C
Measuring range: −200 °C to +1350 °C

## Slide 8: Temperature sensor: Resistance Temperature Detector (RTD)
Principle: Temperature-dependent resistor with well-defined T/R-relation (temp/resistance).
Common: Pt100 RTDs: Pt is platinum. 100 means that R=100 Ohm at 0 degrees C.
R-value is measured with Wheatstones bridge:
[http://en.wikipedia.org/wiki/Resistance_thermometer]

## Slide 9: Select TC or RTD?
Accuracy: RTD is approx. 0.2 deg C. TC is approx. 2 deg C.
In general, is recommended in range -200 - 500 deg C. Above 500 deg C RTD can not be used.
[http://en.wikipedia.org/wiki/Resistance_thermometer]

## Slide 10: Pressure

## Slide 11: Pressure sensor: dP-cell
dP = differential pressure
dP-cells are actually very applicable:
Pressure
Level
Flow

## Slide 12: dP-cell cont.
[Smar Pressure Transmitter.
LD300 Series. Autek AS.]
PH
PL
Pressure is directly applied to the isolating diaphragms (2). Pressure is transmitted to sensing diaphragm through filling fluid (3). Whenever there is difference between PL and PH, the sensing diaphragm (1) , which is a capacitor plate, moves relative to fixed metallized plates (4). The sensing diaphragm (1) deflection results in capacitance variations between the moving and fixed plates. The capacitance variance is detected by electronics.
To electronics
Pressure High
Pressure Low

## Slide 13: dP-cell cont.
Transducer
Transmitter
[Smar Pressure Transmitter. LD300 Series. Autek AS.]

## Slide 14: dP-cell cont.
Application: Boiler control (valve is actuator). 3 dP-cells are used!
[Smar Pressure Transmitter. LD300 Series. Autek AS.]

## Slide 15: Level

## Slide 16: Level sensor: Ultra sound
[Microflex. Autek AS.]
Sound pulses are sent to liquid surface. Refection time indicates level.

## Slide 17: Ultrasound level sensor cont.
[Microflex. Autek AS.]
Characteristics:

## Slide 18: Level measurement with dP-cell

## Slide 19: Flow

## Slide 20: Flow measurement with dP-cell
[http://en.wikipedia.org/wiki/Flow_meter]
From Bernoulli's Law flow and pressure drop across the orifice (measured with a dP-cell) is related as:

## Slide 21: Coriolis
[Heinrichs TMU. Flow-teknikk AS]
Large measurement ranges and high accuracy, e.g.:
0 - 2,2 tons/h with accuracy of 0.1% of measurement span.

## Slide 22: Principle of Coriolis sensor:
Density
Mass flow

## Slide 23: Gas flow sensor: Thermal sensor
[Bronkhorst. Flow-teknikk AS]

## Slide 24: Thermal gas flow sensor cont.:
[Bronkhorst. Flow-teknikk AS]

## Slide 25: Magnetic liquid flow sensor
Induced voltage is proportional to flow velocity, and, therefore, to volumetric flow.

## Slide 26: Doppler (with ultra sound)
Flow is a function of difference between "red" and "blue" travel time.

## Slide 27: Spektroscopic (IR) sensor for gas consentration